[image: image1.jpg]GOOD WILL INSTRUMENT CO.,, LTD.

NO. 7-1, Jhongsing Road, Tucheng Dist., New Taipei City, 236, Taiwan

GE lnerK T (886) 22268-0389 F (886)2 2268-0639  www.gwinstek.com






GDS-122 Specifications

	Conditions for specifications
	The following specifications are applicable when these two conditions are met:

· The GDS-122 has been powered up for at least 30 minutes, during which temperature fluctuation is no more than 5 degrees Celsius.

· The probe attenuation is set to X 10.


Oscilloscope specifications

	Sampling XE "specifications:oscilloscope" 

 XE "oscilloscope:specifications" 
	Mode
	Normal, Peak detection, Average

	
	Rate
	100 MSa/s 

	Input
	Coupling
	DC, AC

	
	Impedance
	1MΩ±2% in parallel with 20pF±3pF 

	
	Probe
	1X, 10X, 100X, 1000X

	
	Max. Input
	400V (peak) 

	
	Channel delay
	150ps (typical)

	Horizontal XE "horizontal:specifications" 
	Sampling rate
	10S/s～100mS/s 

	
	Interpolation
	（sin x）/x

	
	Record length
	6K points on each channel

	
	Scanning speed
	5ns/div～5s/div, 1-2.5-5 step

	
	Sampling rate / relay time accuracy
	±100ppm (time interval ≥ 1ms)

	
	Interval (△T) accuracy (full bandwidth)
	Single: ±(1 interval time +100ppm×reading+0.6ns) Average >16 : ±(1 interval time +100ppm×reading+0.4ns)

	Vertical XE "vertical:specifications" 
	A/D converter
	8 bits resolution (2CH simultaneously)

	
	Sensitivity
	5mV/div～5V/div（at input）

	
	Displacement
	±50V(500mV～5V),±1V(5mV～200mV)

	
	Bandwidth
	20M

	
	Single
	Full bandwidth

	
	Low frequency
	≥5Hz（at input, AD coupling, -3dB）

	
	Rise time
	≤17.5ns (at input, typical)

	
	DC accuracy
	±5% (DC gain)

	
	DC accuracy (avg)
	Avg >16: ±(5% rdg + 0.05 div) for △V

	Trigger XE "trigger:specifications" 
	Sensitivity
	CH1 and CH2: 1div(DC～full bandwidth)

DC coupling: ≥ 50Hz.

	
	Trigger level
	±6 divisions from the screen center

	
	Level accuracy
	±0.3 div (typical, rise/fall time ≥ 20ns)

	
	Displacement
	655div (pre-trigger), 4div (post- trigger)

	
	50% level setting
	Input signal frequency ≥ 50Hz (typical)

	
	Trigger sensitivity
	2 div of peak-to-peak (video trigger)

	
	Signal system
	NTSC, PAL, SECAM (any frequency)

	Measurement
	Cursor
	△V and △T between cursors

	
	Automatic
	Peak-to-peak, average, root mean square, frequency, and cycle

	Probe XE "probe:specifications" 
	
	1X position
	10X position

	
	Bandwidth
	≤ 6 MHz (DC)
	Full bandwidth (DC)

	
	Attenuation rate
	1: 1 
	10: 1

	
	Compensation
	10pf～35pf
	10pf～35pf

	
	Input impedance
	1MΩ±2％
	10MΩ±2％

	
	Input impendence
	85pf～115pf
	14.5pf～17.5pf

	
	Input voltage
	150 V DC
	300V DC


Multimeter specifications

	VDC XE "DCV multimeter:specifications" 

 XE "multimeter:specifications" 

 XE "specifications:multimeter" 
	Input impedance
	10MΩ

	
	Max input
	1000V (DC or AC peak-to-peak value)

	
	Accuracy
	±1%±1 digit

	
	Resolution
	400mV range: 100uV

4V range: 1mV

40V range: 10mV

400V range: 100mV

	VAC XE "ACV multimeter:specifications" 
	Input impedance
	10MΩ

	
	Max input
	750V(AC, virtual value)

	
	Frequency range
	40Hz~400Hz

	
	Display
	Virtual value of sine wave

	
	Accuracy
	±1%±3 digits

	
	Resolution
	4V range: 1mV

40V range: 10mV

400V range: 100mV

	DCA XE "DCA multimeter:specifications" 
	Accuracy
	40mA range: ±1%±1 digit
400mA range: ±1.5%±1 digit
20A range: ±3%±3 digits

	
	Resolution
	40mA range: 10uA

400mA range: 100uA

20A range: 10mA

	ACA XE "ACA multimeter:specifications" 
	Accuracy
	40mA range: ±1.5%±3 digit
400mA range: ±2 %±1 digit
20A range: ±5%±3 digits

	
	Resolution
	40mA range: 10uA

400mA range: 100uA

20A range: 10mA

	Resistance XE "impedance, multimeter:specifications" 
	Accuracy
	400Ω range: ±1%±3 digits

4kΩ~4MΩ range: ±1%±1 digit
40MΩ range: ±1.5%±1 digit

	
	Resolution
	400Ω range: 0.1Ω
4kΩ range: 1Ω
40kΩ range: 10Ω
400kΩ range: 100Ω
4MΩ range: 1kΩ
40MΩ range: 10kΩ

	Capacitance XE "capacitance measurement, multimeter:specifications" 
	Accuracy
	±3%±3 digits

	
	Resolution
	51.2nF range: 10pF

512nF range: 100pF

5.12uF range: 1nF

51.2uF range: 10nF

100uF range: 100nF

	Diode XE "diode measurement, multimeter:specifications" 
	Reading range
	0V~1.5V

	Continuity XE "continuity, multimeter:specifications" 
	Threshold
	< 30Ω


General specifications

	Display XE "display:specifications" 

 XE "specifications:general" 
	Type
	3.8" color liquid crystal display

	
	Resolution
	320 (horizontal) ×240 (vertical) pixels

	
	Color
	4096 colors

	Power XE "power supply:specifications" 
	Consumption
	< 6W

	
	Supply
	100V~240V AC, 50/60Hz

	
	DC input
	8.5VDC, 1500mA

	Environment XE "environment:specifications" 
	Operating
	Temperature: 0  to  40 °C(32  to  104 °F)
Relative humidity: < 75%

	
	Storage
	Temperature: -20  to  60 °C(-4  to  140 °F)
Relative humidity: < 75%

	Mechanical
	Dimension
	18 cm×11.5cm×4cm

	
	Weight
	690g


